## Sheet1
REGULATION / INTERPRETATIONS - REVISIONS
REGULATION # | Issue of Concern
2900-PM-RP003 - Notice to Employees | Should add reference to 25PA 220.9
225.81(a)(3) | Reference to 10cfr34.52 regarding posting - incorrect reference. Should be 10cfr34.53
221.38a c | Does not include the phrase "Upon installation"
221.32a(b)2 | Uses phrase "total misalignment" which can be confusing and lead to an interpretation that both dimensions are added together (L and W) and used to determine compliance.
221.11(b) PA code 49 18.202 | Should the operator of a finger bone densitometer be required to have ARRT cert.
221.11(b) PA code 49 31.31 | Should a Vet assistant be required to have proper training and certification.
???? | All radiation producing machine shall have acceptance testing before being put into service
Electronic Brachytherapy operation should require licensure, similar to accelerator and rad material licensing. Modality is similar to radioactive source brachytherapy, except micro x-ray devices are implanted instead of isotopes.
When we revise the regs we should link 10CFR2206 annual dosimetry reports to open shop x-ray radiography too.
225.102(d) | Shielded room excludes 225.91, 225.92 relating to rad surveys & instrument checks.
227.11a. Equipment requirements. | (a) Open-beam configurations shall have a safety device which either prevents the entry of any portion of an individual’s body into the primary X-ray beam path, or causes the beam to be terminated or interrupted upon entry into the path.
228 | Do we require a written directive for accelerator therapy? It is defined in 215, and referred to in 219 for reportable events, but is it specifically required as in part 35 for RAM?
215.1(e) | What to do about the new NRC security rules. If NRC continues down the road of handling it as a new 10 CFR Part 39 then we will need to implement through license condition until we do a rulemaking and fix 215.1(e). The current incorporation by reference strategy doesn't cover it because we don't incorporate 10 CFR as a whole, we specify the Parts (Chapter equivalents).
228.23a | Change effective range of accelerator licensure from 5 years to 10 years
new addition | mobile x-ray services like visiting nursing homes etc. should be addressed. They are not done on the premises of a doctor or under the direct supervision of a doctor.
223.11(e) | 223.11(e) provides: "(e) The coefficient of variation for exposure may not exceed 0.10 when all technique factors are held constant. This requirement shall be deemed to have been met if, when 4 exposures are made at identical technique factors, the value of the average exposure (a) is greater than or equal to 5 times the maximum exposure (e(max)) minus the minimum exposure (e(min))."
228.72 | applies to units that are capable of both x-ray beams and electron beams. Some units are only capable of electron beams (e.g. tomotherapy) therefore this regulation does it apply.
228.73 | applies to units that are capable of both x-ray beams and electron beams. Some units are only capable of electron beams therefore this regulation does it apply.
221.11(l) | As part of the QA, facilities must have and be able to provide a copy of their operators manual which shall include QA, patient set up, troubleshooting, etc., upon inspection. The reasoning behind this is that all manufacturers have different specs. QA may be different, especially with CR- digital readers with doctors work and acquisition stations. DR - array detectors (Image distance from table may not be indicated on detector which may be found in the owners manual under "specs") with doctors work and acquisition stations as well
221.61(b) | Rule 221.61(b) allows for a more relaxed approach to CT units that are specifically for CT-Sim use. I think the Bureau should consider a similar understanding for CT/PET units. For example, the CT component of our PET/CT system is used EXCLUSIVELY for the purpose of attenuation correction for the pending PET scan. It is not ever used for individual diagnostic CT scans.
221.204(a)(3) | Rule 221.204(a)(3) talks about the need for a QE to perform certain tests "...annually or after any component repair or change which in the opinion of the qualified expert may effect the performance of the CT unit…". I have always interpreted the fact that there is no specific time frame linked to this rule as an understanding on the part of the DEP that there simply aren't enough physicists in the Commonwealth to meet the demands of a "…before any patients can be scanned" kind of provision. Personally, I try to get there as soon as possible after, say, a tube change, but have told facilities that doesn't mean they have to shut down all patient exams.
225.102(d) | the shielded room regs still, by exempting portions of 225, do not require: calibration and function tests of survey meters; tests of interlocks; records of survey or interlock checks.
223.22 | Requires veterinarians to comply with 10 CFR 35 Subparts F, G, H, and K for use of sealed sources. This far exceeds current NRC licensing practice and guidance, which regulates the veterinary use of RM under 10 CFR Parts 30 and 31. See NUREG-1556 v7 App. H for licensing guidance. See http://www.nrc.gov/materials/miau/vet-toolkit.html
223 | No provision for the regulation of unsealed sources in veterinary medicine. Add a section (223.23?) to regulate the use of unsealed sources in veterinary medicine under 10 CFR 30 and 31.
217.137 | De-regulate Re-187 - ALI (600,000 μCi) equivalent to 12.8 MT of material
217.131(b) | 10 CFR 30.41(a)(6) does not exist, but 30.41(b)(6) does) - check
216.1 | Notes for EBX - 216.1 After(b) Need(c) exempt needed from 216.2EBX= License relating to as defined in 221.1
216.3 | 216.3 Exemption Need after (4) to put in (5) EBX shall be Licensed
218.11 | 218.11After (d) need to enter/ EBX License fee$1000
221.1 | 221.1 Word change needed at this sentence(A registrant/ License who uses x-rays in the healing)
221.2 | 221.2. Definitions will be needed? Need to provide / or should this go in chapter equipment requirements 221.71. (therapeutic x-ray system) "Electronic Brachytherapy" means a method of radiation therapy where an electrically generated source of ionizing radiation is placed in or near the tumor or target tissue to deliver therapeutic radiation dosage.
221.71 | 221.71. Equipment requirements. Filters, Tube housing (ebx) must have an exclusion from this area
§221.204 (a)(3)(i) | 25 PA Code §221.204 (a)(3)(i) relative to the measurement of beam quality for computed tomography units is not a concern for patient radiation dose or patient radiation safety, because all manufacturers add heavy aluminum beam shaping filters which provide a high degree of low energy beam filtration. Furthermore, beam attenuation is measured throughout the patient scan and incorporated into the reconstruction of the images. If there were any problems with beam quality, it would be immediately noticeable in patient images and expected Hounsfield units. The State of Pennsylvania is the only state to have written this requirement into their regulations in 1998. Testing of this performance parameter is not component of any national or international physics testing standards for computed tomography systems.
§221.204 (a)(5) | 25 PA Code §221.204 (a)(5) relative to determination of mR/mAs for head and body is a redundant requirement. Measurements are already required be performed of typical doses to head and body with head and body dose phantoms (25 PA Code §221.204(a)(1)). The resultant computed tomography dose index (CTDI) is compared to the manufacturer’s FDA submission for this parameter, and measured to assess the clinical protocols used and the site for comparison to various national and international reference levels. Inherent in the measurements made in the 2 dose phantoms is an mR/mAs value. Therefore the additional requirement to calculate an mR/mAs value in head and body dose phantoms is of no clinical or patient safety value.
219.3 | Clarify definition for MRE, particularly (ii)
228.21a (c) and (d) | delete both sections - they were initially recorded to account for 'grandfathered' accelerators. We have no more 'grandfathered' accelerators. All accelerators are accounted and licensed.
219.6(7) | Change this to reflect that radiation exposure reports required by 10 CFR 20.2206 be submitted to the REIRS program as specified in 10 CFR 20.1007 and are not to be reported to the DEP.
228.72 & 228.73 | These regulations refer to a record/verify system. The license should include a special condition to designate the record/verify system used per unit.
221.25 Table 2 | update to federal regs 21-CFR-1020.30
Non-medical cabinet CT devices should be inspected per the Non-Med Cabinet regulations
225.4a | reference 10cfr2206
Chapter 223 | should include "Registrant's Responsibilities" similar to 221.11
Chapter 223 | should address CT's
217.2(c ) | delete
219.5(b) | delete (a)(1) and leave it as 20.2206
219.6(7) | delete "and, for NRC licenses, to the NRC until agreement state status is in effect"
218 Appendix A footnote 1
Lately, we have been getting a few requests for "reciprocity" based on GL's in other states, I have explained that we cant issue a reciprocity on a GL, but here is the problem that they are wondering..why cant they just get a GL in PA with only Temp. job sites and we give them the 72 hr. condition? Should we change the regs to try and stop this somehow?
Chapter 223 | Add "Registrant Responsibilities" to veterinary chapter similar to 221.11. Include 221.11(a, b, d, e, l, m)
221.2 | delete the first of two "Certified Components" definitions
221.2 | Tighten up the healing arts screening definition to include facilities like XXX Hospital – perhaps include advertising or targeting asymptomatic members of the public in the definition.
Require X-Ray operators in healing arts facilities (find term to include hospitals and other medical facilities but exclude dentists and other specialists) to be an RTs (Radiation Technologists) under DOS, ARRT, ISCD or another recognized board
Register physicists (medical and health physicists) to be qualified experts in PA
218.12 | This section should be revised to add license modification and/or amendment to enforce against licensees with delinquent fees and sealing off sources to enforce against registrants with delinquent fees.
mobile x-ray services like visiting nursing homes etc. should be addressed. They are not done on the premises of a doctor or under the direct supervision of a doctor. The regs should formally recognize our long standing approval of that special category of registrants and list any restrictions.
§ 218.11 (c) | This section should be revised to indicate that annual license fees shall be paid according to the fee schedule set forth in Appendix A. (We spend time indicating that the Department will send an invoice at least 2 mo. prior to the license expiration. However, many licenses extend for 10 years. This creates a scenario where a license continues to be in place, but they are not specifically compelled to pay the annual fee.) Later in § 218.12, we say that they can be subject to penalties for non-payment. However, the regs again fail to say that a licensee shall pay. We need to be clearer on this topic because the current lack of clarity creates issues from an enforcement perspective.
§ 215.2 | Could we add a proper definition of a qualified individual/qualified expert? It would be helpful to include a statement pertaining to training received or testing that a person has underwent. Our regs reference these terms but there is no standard to hold them to.
221.41a | The amount of fluoro time should be recorded somewhere, preferably the patient chart.
225.102(d) | I Read you response in line 16 ' initial shielding evaluation should be sufficient.' I think the meter should at least be function checked. Also, if cabinet systems are to be surveyed after maintenance and interlocks checked routinely, the same requirements should be for shielded rooms, which typically have higher energy x-rays.
227.21 thru 23; 227.31-33; 227.101 -104 | Are x-ray gauging equipment, e-microscopes, and calibration system exempt from the rest of 227? Like the requirements for annual surveys etc. If so, it should be noted.
228.33a(2) | If the material is not moving, should they be required to observe it? If the intent is to monitor the room to make sure nobody is in it (as Ray had suggested), then it should be stated that way to require the viewing system to 'see' the entire room.
§ 221.30 | Can this be amended to include certified systems, as well? At times certified systems are found to have a COV exceeding 0.10.
221.2 | RPAC member, wants a better definition for “Licensed Practitioner of the Healing Arts” .
218 Appendix A | Need a fee category for Generally Licensed source material covered by 10 CFR 40.25. Revise 3Q to also refer to 217.171
216.2b | subsection (a)(1) should non-medical vendors be allowed 30 days or quarterly notification, rather then 15 days per FDA requirement?
217/143 | Need to catch Cd-109 in the GL registration now that its no longer accelerator produced (31.5 does not cover it)
215.1 | need to include 30.32(e) which points to part 170 for fee schedule, in list of "not incorporated" parts
230.13 | This paragraph appears to say any licensed material needs to meet DOT requirements even if it is below DOT RAM limits. The paragraph needs to include how it shall be shipped, i.e.. Lim quantity, LSA, type A or what.
228 | SEPTA and some well loggers possess a portable neutron generator described in SS&D CO-1012-D-101-S for detection of explosives and well logging. This is a particle accelerator operating with a 100-150 kV potential (note this increased kV is a change to the SS&D). This accelerator should be controlled under 228. Currently these units are controlled under RAM license for H-3, this does not address the risk posed by this unit. Adding "explosives detectors and well logging" to paragraph 228.45 may fix this issue even if the units are not licensed as accelerators.
228.35(g) | 228.35(g)(5) should not be limited to medical use, instead it should be separated to 228.35(h)
216.2b | include registrant's responsibility for radiation protection, in particular, reference 219.5 requirements for badging/monitoring employees
218.11(i) | We're required to review fees and provide a report to EQB. It should record something like: “the fees have been reviewed and are considered adequate at this time”
228.35( c) | Testing scram buttons (cutoff switches) every 3 months is hard on accelerators. The frequency for Medical accelerators should be decreased to 6 months. Non-Medical accelerators should stay at 3 months.
228 | We should have a requirement to have accelerator Vaults room name/number Placard onto the vault door or door frame to match the application room number. This way any inspector can use the License in hand and walk from Vault to Vault identifying each accelerator from the license to the assigned Vault Room placard. This will help reduce confusions with same manufactured accelerators that are set up with different modalities and energy’s. As of now you would have to take the accelerator cover off the accelerator and look for the serial number to identify each machine
221.62(?) | Stereotactic Breast Biopsy Machines - Need to get these machines registered ("Other Systems") and set an inspection protocol for these devices (similar to mammo units but not as involved or review physicist's report)?
218.11a | delete entire section
227 | Address CT devices in a non-medical application. Should fall under the "analytical" section
Include (possibly by reference) a mechanism to 'ban' a qualified person who shows bad or inappropriate judgment in performing his duties.
230.13 | 230.13 appears to say that any shipment of licensed material is to be controlled as a radioactive shipment. But 49 CFR 171.2(k) makes it illegal to ship material as hazardous (radioactive) unless it meets DOT's definition of radioactive. Licensed material that is not allowed for Decay-In-Storage is always licensed but may not require DOT hazardous controls for shipment.
218.11 | I just spoke to Ken and he suggested we may want to consider adding an application fee when we revise the regs, since approving/denying an application costs the Department time and money and we are a self-funded program.
216.7 | Change from Out-of State radiation producing machines TO: Temporary use of radiation producing machines... we have in state vendors doing demo's also
Let's make sure we get a BRP reg rev drafted for requiring reporting to BRP any tech / natural hazards impacting RAM safety or security... e.g., earthquakes, floods, power outage, etc.
221.21 | Suggest that ANY human use component or assembly require compliance w/ FDA regs (so as not to allow use of non-certified items from both domestic and international suppliers)
216.7 | add all ‘mobile’ registrants notify us when and where they will be operating within our Commonwealth, similar to 217.203(a)(2).
In addition to the growing use of X-ray machines for blood irradiation, there are also some facilities who are performing other (package) sterilization, and the design of these facilities, the X-ray devices, and operator training and protection ought to be addressed in some sort of regulatory development.
221.3 | Amend to include certified systems. Certified systems are being found with fluctuating exposure rates with COV exceeding 0.10 (up to 30%).
215.32 | include a statement NOT allowing an exemption to our fee requirement
223 | include something regarding protection of animals during research. Perhaps reference USDA's Animal Welfare Act and FDA/NRC policy regarding humane care and use of laboratory animals
Transportation section - add some kind of hook requiring DOT exemption
TENORM may eventually require licensing - check out CRCPD's SSR
215.31 | Add to the end of this section: "However, the department may not exempt or waive the payment of any fee required under this article."
225? | Add a section to address c-arm use for checking tires and other industrial applications
219 | standardize medical event definition for X-rays and materials
227 | standardize X-ray and materials portable XRF unit requirements
include recommendations for digital imaging
216.3 (1) states that electron bean welders (and melters I suspect) are not exempt from registrations. There is no mention of EBs anywhere so are they under 225 or 227 with NEMICs?
240 | Form 3 is specific to licensees and registrants and not mentioned in the radon world. It should be referenced because there are testing and mitigation firms as well as labs and chambers that have employees.
221.11 (l) | Suggest changing QA related records retention from three (3) years to five (5) years.
216.2b (1) | The following information shall be submitted ( in writing) .. Should be replaced with utilizing the PA Vendor/Service Providers Report to BRP form. (2900-PM-RP0046)
216.2b (2) | Remove this section. Utilizing the Pa report Form will elevate having to identify registrants with inaccurate address information help identify who needs to be registered.
228.75 (6) | Remove "end effect" This was for Co 60 time it took for source to get in and out of placement..
230.15 (new) | add new section '230.15' titled: "Transportation of unlicensed material" Material not licensed by the department or under the specific regulatory control of another state or federal agency, but meeting the definition of radioactive material in the United States Department of Transportation regulations in 49CFR, shall conform to the standards and requirements of those regulations.
20.2206 | address reporting requirements of individual monitoring
230.13 | Expand to include licensee "and non-licensee"