## Instructions
Introduction
This spreadsheet can be used to calculate and document the analyst Initial Demonstration of Capability (IDC) for the grab sample method for chlorine analysis, required by EPA Method 334.0. It consists of 5 tabs: Instructions, Blank IDC form, IDC example - PASS, IDC Example - FAIL, and Dilutions.
The analyst IDC assesses the accuracy and precision of the analytical technique. Per Method 334.0, it must be completed by each person who conducts chlorine analysis prior to using the grab method for compliance monitoring or to verify the accuracy of on-line chlorine analyzers used for compliance monitoring. The IDC consists of three steps: analysis of a method blank to determine the degree of interference by the reagent, analysis of five replicates of a single concentration of a prepared primary standard, and calculations to determine the initial demonstration of accuracy and initial demonstration of precision.
Method Blank: For the method blank, the analyst uses reagent grade water (ideally organic-free or chlorine-demand-free water) as the sample, following all steps in the chosen grab sample method to complete the analysis. Any measured chlorine concentration is attributable to interference from the reagent.
Five Replicates of Primary Standard: The analyst must prepare a diluted primary standard, at a concentration that falls within the range of the selected grab sample method. (See the Dilutions tab of this spreadsheet for more information.) Then, the analyst measures the concentration of five samples of the prepared standard, treating each as a discreet sample.
Initial Demonstrations of Accuracy and Precision: These are completed through a series of calculations. For the initial demonstration of accuracy, the percent difference of the average of the five replicates is calculated. Percent difference must be ±15% of the reference sample concentration. For the initial demonstration of precision, the relative standard deviation of the five replicates is calculated, and must be ≤15%. The Blank IDC Form tab of this spreadsheet is set up to complete these calculations. If both acceptance criteria are met, the analyst has passed the IDC and can begin to use the grab method for compliance measurements.
Blank IDC Form tab
Use this tab to record all data for the analyst IDC.
IMPORTANT: Only record data in the blue shaded cells.
The orange shaded cells contain formulas for calculations and should not be edited.
Part I. General Information:
Record the following information in the blue shaded cells: the date the IDC is conducted, system PWSID, system name, analyst name, analysis method, and analyzer manufacturer, model, and serial number. Note that each person who conducts chlorine analysis must perform an Initial Demonstration of Capability prior to using the grab sample method for compliance monitoring or to verify the accuracy of on-line chlorine analyzers used for compliance monitoring.
Part II. Independent Reference Sample Analyses:
Use this section to record the Independent Reference Sample Concentration (A), which is the concentration of the diluted primary standard prepared prior to analysis. Also use the table in this section to record the measured concentrations of the method blank and the five replicates of primary standard. The corrected concentrations (B) will be calculated by subtracting the method blank from each of the five measured concentrations, and the average of corrected concentrations (C) will be calculated by dividing the sum of the corrected concentrations by five.
Part III. Initial Demonstration of Accuracy:
In this section, the percent difference will be calculated. Percent difference must be within ±15% of the reference sample concentration. If this criterion is met, the calculated percent difference will be displayed in green, and the answer to the question "Is the average concentration of the 5 replicates (C) within ±15% of the independent reference sample concentration (A)?" will be "YES", which will also be displayed in green. If the criterion is not met, the calculated percent difference will be displayed in red, and the answer to the question will be "NO" and will also be displayed in red. Note: this section contains no blue shaded cells to fill in, only calculations.
Part IV. Initial Demonstration of Precision:
In this section, the standard deviation (S) and relative standard deviation (RSD) will be calculated. The relative standard deviation must be ≤15%. If this criterion is met, the calculated relative standard deviation will be displayed in green and the answer to the question "Is the Relative Standard Deviation (RSD) ≤15%?" will be "YES", which will also be displayed in green. If the criterion is not met, the calculated relative standard deviation will be displayed in red, and the answer to the question will be "NO" and will also be displayed in red. Note: this section contains no blue shaded cells to fill in, only calculations.
Calibration Standard Solution:
Record important information on the primary standard solution used to complete the IDC. Fill in the blue shaded cells with the primary standard manufacturer, lot number, and expiration date.
Initial Demonstration of Capability: PASS/FAIL
If the criteria for the demonstrations of accuracy and precision are met, "PASS" will be displayed in green. If the criteria are not met, "FAIL" will be displayed in red.
It is important to note that if the analyst does not pass the IDC attempt, the procedure should be evaluated so the source of error can be determined and addressed. Once any issues have been identified, the analyst should repeat the IDC procedure. If the analyst is not able to pass the IDC, he or she should not conduct compliance measurements for chlorine residual.
If the analyst successfully passes the IDC, he or she should print a copy of the completed spreadsheet to retain for documentation.
IDC Example - PASS tab
This tab contains an example of a completed IDC that the analyst has successfully passed.
IDC Example - FAIL tab
This tab contains an example of a completed IDC that the analyst has NOT passed, because the acceptance criteria have not been met.
Dilutions tab
This tab can be used in planning and calculating the diluted primary standard for the IDC.
Use the blue shaded cells to input data.
The orange shaded cells contain formulas for calculations and should not be edited.

## Blank IDC Form
EPA METHOD 334.0 RECORD OF
ANALYST INITIAL DEMONSTRATION OF CAPABILITY
I. General Information
Date:
| PWSID: |  |  |  | System Name:
| Analyst Name1: |  |  |  | Analysis Method:
| Analyzer |  |  | Analyzer |  | Serial
| Manufacturer: |  |  | Model: |  | Number:
1 | Each person who conducts chlorine analysis must perform an Initial Demonstration of Capability prior to using the grab sample method for compliance monitoring or to verify the accuracy of on-line chlorine analyzers used for compliance monitoring.
II. Independent Reference Sample Analyses
| Independent Reference Sample Concentration (A): | [5 empty] | mg/L
Note: For details on calculating and preparing a diluted standard, please see the Dilutions tab or job aid.
| SAMPLE | Measured Concentration (mg/L) |  | Corrected Concentration (B) Measured Conc. - Method Blank (mg/L)
Method Blank 2
| 1 |  |  | 0
| 2 |  |  | 0
| 3 |  |  | 0
| 4 |  |  | 0
| 5 |  |  | 0
| Average of Corrected Concentrations (C): |  |  | 0
2 | Method blank must be ≤1/3 the concentration of the lowest standard used to verify the calibration curve.
III. Initial Demonstration of Accuracy
#DIV/0!
| Percent Difference = |  | ((C - A) / A) x 100% = |  |  | %
[6 empty] | Is the average concentration of the 5 replicates (C) within ±15% | #DIV/0!
of the independent reference sample concentration (A)?
IV. Initial Demonstration of Precision
Standard Deviation (S):
0
| Standard Deviation = |  | √(Average of (B-C)2) =
Relative Standard Deviation (RSD):
#DIV/0!
| Relative Standard Deviation = |  |  | (S/C)x100 =
#DIV/0!
Is the Relative Standard Deviation (RSD) ≤15%?
Calibration Standard Solution:
Primary Standard Manufacturer:
Expiration Date:
Lot Number:
Initial Demonstration of Capability: |  |  |  |  | PASS/FAIL |  | #DIV/0!
In order to pass the Initial Demonstraton of Capability:
1. Percent Difference must be within ±15% for the Initial Demonstration of Accuracy, and
2. Relative Standard Deviation must be ≤15% for the Initial Demonstration of Precision.
NOTE: If the analyst is not able to pass the IDC, the procedure should be evaluated
so the source of error can be determined and addressed. Once any issues have been
identified, the analyst should repeat the IDC procedure.
IF THE ANALYST IS NOT ABLE TO PASS THE IDC, HE OR SHE SHOULD NOT CONDUCT
COMPLIANCE MEASUREMENTS FOR CHLORINE RESIDUAL!
If the analyst successfully passes the IDC, it is important to print and maintain a copy
of this form as documentation.

## IDC Example - PASS
EPA METHOD 334.0 RECORD OF
ANALYST INITIAL DEMONSTRATION OF CAPABILITY
I. General Information
[6 empty] | Date: | Sept. 10, 2018
| PWSID: | 1234567 |  |  | System Name: | ABC Water Company
| Analyst Name1: | Joe Operator |  |  | Analysis Method: | DPD
| Analyzer |  |  | Analyzer |  | Serial
| Manufacturer: | Hach |  | Model: | PC II | Number: | 123456789XYZ
1 | Each person who conducts chlorine analysis must perform an Initial Demonstration of Capability prior to using the grab sample method for compliance monitoring or to verify the accuracy of on-line chlorine analyzers used for compliance monitoring.
II. Independent Reference Sample Analyses
1
| Independent Reference Sample Concentration (A): | [5 empty] | mg/L
Note: For details on calculating and preparing a diluted standard, please see the Dilutions tab or job aid.
| SAMPLE | Measured Concentration (mg/L) |  | Corrected Concentration (B) Measured Conc. - Method Blank (mg/L)
| Method Blank 2 | 0.01
| 1 | 0.9 |  | 0.89
| 2 | 0.87 |  | 0.86
| 3 | 0.93 |  | 0.92
| 4 | 0.84 |  | 0.83
| 5 | 0.87 |  | 0.86
| Average of Corrected Concentrations (C): |  |  | 0.872
2 | Method blank must be ≤1/3 the concentration of the lowest standard used to verify the calibration curve.
III. Initial Demonstration of Accuracy
-12.8
| Percent Difference = |  | ((C - A) / A) x 100% = |  |  | %
[6 empty] | Is the average concentration of the 5 replicates (C) within ±15% | YES
of the independent reference sample concentration (A)?
IV. Initial Demonstration of Precision
Standard Deviation (S):
0.030594
| Standard Deviation = |  | √(Average of (B-C)2) =
Relative Standard Deviation (RSD):
3.5085
| Relative Standard Deviation = |  |  | (S/C)x100 =
YES
Is the Relative Standard Deviation (RSD) ≤15%?
Calibration Standard Solution:
|  |  | Primary Standard Manufacturer: | Hach
|  |  | Expiration Date: | Oct. 2020
|  |  | Lot Number: | A9999
Initial Demonstration of Capability: |  |  |  |  | PASS/FAIL |  | PASS
In order to pass the Initial Demonstraton of Capability:
1. Percent Difference must be within ±15% for the Initial Demonstration of Accuracy, and
2. Relative Standard Deviation must be ≤15% for the Initial Demonstration of Precision.
NOTE: If the analyst is not able to pass the IDC, the procedure should be evaluated
so the source of error can be determined and addressed. Once any issues have been
identified, the analyst should repeat the IDC procedure.
IF THE ANALYST IS NOT ABLE TO PASS THE IDC, HE OR SHE SHOULD NOT CONDUCT
COMPLIANCE MEASUREMENTS FOR CHLORINE RESIDUAL!
If the analyst successfully passes the IDC, it is important to print and maintain a copy
of this form as documentation.

## IDC Example - FAIL
EPA METHOD 334.0 RECORD OF
ANALYST INITIAL DEMONSTRATION OF CAPABILITY
I. General Information
[6 empty] | Date: | Sept. 10, 2018
| PWSID: | 1234567 |  |  | System Name: | ABC Water Company
| Analyst Name1: | Joe Operator |  |  | Analysis Method: | DPD
| Analyzer |  |  | Analyzer |  | Serial
| Manufacturer: | Hach |  | Model: | PC II | Number: | 123456789XYZ
1 | Each person who conducts chlorine analysis must perform an Initial Demonstration of Capability prior to using the grab sample method for compliance monitoring or to verify the accuracy of on-line chlorine analyzers used for compliance monitoring.
II. Independent Reference Sample Analyses
1
| Independent Reference Sample Concentration (A): | [5 empty] | mg/L
Note: For details on calculating and preparing a diluted standard, please see the Dilutions tab or job aid.
| SAMPLE | Measured Concentration (mg/L) |  | Corrected Concentration (B) Measured Conc. - Method Blank (mg/L)
| Method Blank 2 | 0.01
| 1 | 0.75 |  | 0.74
| 2 | 0.87 |  | 0.86
| 3 | 0.93 |  | 0.92
| 4 | 0.84 |  | 0.83
| 5 | 0.87 |  | 0.86
| Average of Corrected Concentrations (C): |  |  | 0.842
2 | Method blank must be ≤1/3 the concentration of the lowest standard used to verify the calibration curve.
III. Initial Demonstration of Accuracy
-15.8
| Percent Difference = |  | ((C - A) / A) x 100% = |  |  | %
[6 empty] | Is the average concentration of the 5 replicates (C) within ±15% | NO
of the independent reference sample concentration (A)?
IV. Initial Demonstration of Precision
Standard Deviation (S):
0.058788
| Standard Deviation = |  | √(Average of (B-C)2) =
Relative Standard Deviation (RSD):
6.9819
| Relative Standard Deviation = |  |  | (S/C)x100 =
YES
Is the Relative Standard Deviation (RSD) ≤15%?
Calibration Standard Solution:
|  |  | Primary Standard Manufacturer: | Hach
|  |  | Expiration Date: | Oct. 2020
|  |  | Lot Number: | A9999
Initial Demonstration of Capability: |  |  |  |  | PASS/FAIL |  | FAIL
In order to pass the Initial Demonstraton of Capability:
1. Percent Difference must be within ±15% for the Initial Demonstration of Accuracy, and
2. Relative Standard Deviation must be ≤15% for the Initial Demonstration of Precision.
NOTE: If the analyst is not able to pass the IDC, the procedure should be evaluated
so the source of error can be determined and addressed. Once any issues have been
identified, the analyst should repeat the IDC procedure.
IF THE ANALYST IS NOT ABLE TO PASS THE IDC, HE OR SHE SHOULD NOT CONDUCT
COMPLIANCE MEASUREMENTS FOR CHLORINE RESIDUAL!
If the analyst successfully passes the IDC, it is important to print and maintain a copy
of this form as documentation.

## Dilutions
To calculate the required volume of certified primary standard stock solution needed to mix the desired volume of check standard,
use the dilution equation C1V1 = C2V2:
C2: | 1 | mg/L | Desired concentration for calibration check standard to prepare
V2: | 100 | mL | Desired volume of prepared check standard (i.e. volume of class A volumetric flask)
C1: | 66.5 | mg/L | Manufacturer's certified concentration of primary standard
V1: | 1.5038 | mL | Calculated volume of primary standard to mix desired volume (V2) of check standard
Use a pipette to measure the calculated V1 volume into a class A volumetric flask that matches the V2 volume
and dilute it to the calibration mark.
To calculate the exact concentration of a diluted check standard using a known volume of primary standard stock solution,
rearrange the above formula and solve for C2:
C1: | 66.5 | mg/L | Manufacturer's certified concentration of primary standard
V1: | 1 | mL | Volume of primary standard stock solution used to prepare diluted check standard
V2: | 100 | mL | Desired volume of prepared check standard (i.e. volume of class A volumetric flask)
C2: | 0.665 | mg/L | Calculated concentration of diluted check standard
The Department has additional resources for assistance on calculating and preparing diluted check standards on the
Method 334.0 website, located at the following link:
https://www.dep.pa.gov/Business/Water/BureauSafeDrinkingWater/DrinkingWaterMgmt/Pages/Method-334.0.aspx
Tutorial 3 in the video series on Method 334.0 contains a demonstration on the calculation and preparation of a diluted standard.
There is also a Job Aid on mixing Primary Standards.